## Slide 1: Kwaliteitscontrole in de x-stralenbeeldvorming van de borst
Hilde Bosmans & het mammo (QC) team

## Slide 2: Overzicht
Set the scene
De dosis in de mammografie
15 jaar dagelijkse QC en fysisch technische QC in borstkankerscreening
Basisprincipes van nieuwe technieken in de borstbeeldvorming
  DBT
  Synthetic 2D
QC in DBT & synthetische 2D
Samenvatting

## Slide 3: Set the scene

## Slide 4: Set the scene

## Slide 5: Set the scene

## Slide 6: Set the scene

## Slide 7: Set the scene
benefit

## Slide 8: Er is kwaliteitscontrole in de Vlaamse (Belgische) screening
Wat kunnen we zelf nog doen???

## Slide 9: Risico’s gerelateerd aan screening
De ‘Mean Glandular Dose’, en de kans op inductie van kanker tgv de straling
Schrik of stress in het algemeen
Slechte beeldkwaliteit (missen van letsels)
Slechte positionering
Verouderde behandelingswijzen
Overdiagnosis, over treatment
Vals gevoel van veiligheid
…

## Slide 10: De dosis in de mammografie

## Slide 11: MGD = K * ……
K, ‘incident air kerma’, de hoeveelheid straling gemeten onder de compressieplaat, waar de straling in de borst gaat,
Wordt bepaald door anode/filter, de mAs, de kV, de dikte van de borst
Add. to the
European Guidelines
De Mean Glandular Dose (MGD)
Court. N Geeraert

## Slide 12: MGD = K * g ……
g = conversie naar een 50% vet, 50% glandulaire borst met 2 x 0.5cm huid
De Mean Glandular Dose (MGD)

## Slide 13: MGD = K * g * c……
c = aanpassing aan de glandulariteit van de borst van de
patiënt
Fatty
Dense
De Mean Glandular Dose (MGD)

## Slide 14: MGD = K * g * c……
D Dance, 2000
Estimates of average breast composition for different compressed breast thickness in Guildford (open circles) and the Northern Region (full circles), for women in the age range 50 to 64y.
De Mean Glandular Dose (MGD)

## Slide 15: MGD = K * g * c * s
s = compensatie voor het X-stralen spectrum, meer bepaald, de materialen van anode en filter (vroeger: Molybdenium-Molybdenium, nu ook andere anode-filter materialen)
Mo/Rh
W/Ag
W/Al
Rh/Rh
De Mean Glandular Dose (MGD)
Rh/Ag
W/Rh

## Slide 16: MGD = K * g * c * t ……
Projection angle (degrees)
Tomo factor t
All spectra, thicknesses and glandularities
Courtesy K Young
De Mean Glandular Dose (MGD) (bij DBT)

## Slide 17: Welke dosislimieten zijn er voor de vrouw ?
Geen !

## Slide 18: Er zijn wel limieten…
die we controleren met testobjecten
Testobjecten zijn representatief voor de ‘gemiddelde’ borst
  Dikte testobject van 2cm -> 7cm, i.e. borst van 2cm -> 9cm

## Slide 19: Er zijn wel limieten…

## Slide 20
Binnen deze limieten moet
een kwaliteitsnorm
gehaald wordenen !

## Slide 21
Belangrijkste Criterium: Hoe dik moet een kleine goudstip zijn zodat men die nog kan zien in een beeld?
‘CDMAM’ test object = contrast –detail test object voor de mammografie
Technische beeldkwaliteit in de mammografie

## Slide 22
Een typisch resultaat:

## Slide 23
Technische beeldkwaliteit in de mammografie

## Slide 24
Indien nodig en mogelijk, wordt de technische beeldkwaliteit in de mammografie aangepast

## Slide 25
+ Tube output (K)
+ beam quality ( s )
Age or glandularity ( c)
-> 3-jaarlijkse Dosis steekproef (FANC)

## Slide 26
Typisch voorbeeld van datacollectie

## Slide 27: De nationale DRLs

## Slide 28: Dosismonitoring & management in vgl met DRLs

## Slide 29: Dosismonitoring & management in vgl met DRLs
Leuven QC data

## Slide 30: ‘Mean glandular dose’ in de Vlaamse screening

## Slide 31: Dosislimieten
Nooit voor individuele patiënten
European Guidelines: ‘acceptable’ en ‘achievable’ levels voor 2D mammo
Gewoonlijk worden dezelfde niveaus aangehouden voor tomosynthese
Uitgewerkte Nationale Diagnostic Reference Levels (DRL)

## Slide 32: Fysisch-technische kwaliteitscontrole in de screening
Type test
Acceptatie test
  (half)jaarlijkse test
Dagelijkse test

## Slide 33: Dagelijkse kwaliteitstesten
Dagelijkse test = 2 opnamen van een homogene plaat
Idee: een opname van een homogene plaat, zou een homogeen beeld moeten opleveren.
Voorbeeld van calibratie artefact
Qaelum NV
www.qaelum.com

## Slide 34: Voorbeeld
Stofje op testobject of artefact?
Qaelum NV
www.qaelum.com

## Slide 35: Dagelijkse kwaliteitstesten
Dode detector (halve) lijn
Qaelum NV
www.qaelum.com

## Slide 36: Dagelijkse kwaliteitstesten
Monitoren,dagelijks getest met variabel patroon -> ‘goede kijkomstandigheden’
Qaelum NV
www.qaelum.com

## Slide 37: Welkom op onze voorstelling op de IWBI 2020

## Slide 38: Introduction
More than 15 years ago we started with daily QC supervision in digital mammography
Later digital systems have become more robust
Is daily QC still of benefit for modern units?
This study: Overview of problems encountered with daily QC during the period January 2015 – January 2020 + 5 year trend analysis

## Slide 39: The network of digital mammography systems
January 2015 102 devices (86 digital)
↓
January 2020 107 devices (105 digital)
Systems are operated in Belgian breast screening programmes
All major vendors represented
Chart: 
| Fuji Amulet | Fuji Amulet f | Fuji Amulet S | Fuji Amulet Innovality | GE Senographe DS | GE Senographe Essential | GE Senographe Pristina | Hologic Lorad Selenia | Hologic Selenia Dimensions | Hologic 3Dimensions | IMS Giotto Class | IMS Giotto Rafaello | Sectra Microdose | Philips L30 MDM | Philips Microdose SI L50 | Planmed Clarity | Siemens Mammomat Novation | Siemens Mammomat Inspiration | Siemens Mammomat Fusion | Siemens Inspiration PRIME | Siemens Mammomat Revelation | CR | Film-screen
January 2015 | 2 | 0 | 4 | 3 | 2 | 14 | 0 | 9 | 12 | 0 | 0 | 3 | 6 | 2 | 0 | 0 | 1 | 26 | 0 | 2 | 0 | 13 | 3
January 2020 | 1 | 1 | 3 | 10 | 1 | 15 | 3 | 7 | 10 | 6 | 1 | 6 | 4 | 1 | 1 | 1 | 0 | 14 | 3 | 11 | 6 | 1 | 1

## Slide 40: Remotely controlled daily Quality Control
Performed by local radiographers
2 ‘For Processing’ images of a 4cm PMMA block
  Under AEC
  Geometry used for breast cancer screening
  With grid
  Using rigid compression plate
  180˚ rotation of PMMA between 2 exposures
Data send to the physicists for analysis
Using commercially available software (Qaelum NV, Leuven, Belgium)

## Slide 41
Data
Date and time of received data
DICOM header tags
Thumbnail images: mean pixel value signal to noise ratio peak variance
Noise power spectra
Part of original image automatically cropped around localized artefacts

## Slide 42: Data
Date and time of received data
DICOM header tags: some always visualized others visible via input panel
Thumbnail images: mean pixel value, signal to noise ratio, variance
Part of original image automatically cropped around localized artefacts
Noise power spectra (currently not used)

## Slide 43: Results: problems encountered Jan ‘15 – Jan ‘20
 | Pixels |  | Lines |  | Blocks |  | Ghosts |  | Reprod |  | Inhomogeneities |  | Collimators |  | Other |  | Total
2015 | 12 | 46% | 3 | 12% | 1 | 4% | 4 | 15% | 2 | 8% | 4 | 15% |  | 0% |  | 0% | 26
2016 | 19 | 56% | 5 | 15% | 1 | 3% |  | 0% | 4 | 12% | 3 | 9% | 1 | 3% | 1 | 3% | 34
2017 | 42 | 54% | 8 | 10% | 5 | 6% | 3 | 4% | 11 | 14% | 5 | 6% | 4 | 5% |  | 0% | 78
2018 | 25 | 64% | 3 | 8% | 5 | 13% | 3 | 8% | 1 | 3% | 2 | 5% |  | 0% |  | 0% | 39
2019 | 40 | 49% | 9 | 11% | 2 | 2% | 4+24 | 34% |  | 0% | 1 | 1% | 2 | 2% |  | 0% | 82
Total | 138 | 53% | 28 | 11% | 14 | 5% | 14+24 | 5% + 9% | 18 | 7% | 15 | 6% | 7 | 3% | 1 | 0% | 259
All artefacts subdivided in 8 categories
Pixels: localized artefacts
Lines: line or line segment artefacts
Blocks: special pattern or detector blocks visible
Ghosts: lag and ghosts
Reprod: sudden changes in exposure signifiers
Inhomogeneities
Collimators: collimator problems
Other
Chart: 
| Pixels | Lines | Blocks | Ghosts | Reprod | Inhomogeneities | Collimators | Other
2015 | 12 | 3 | 1 | 4 | 2 | 4
2016 | 19 | 5 | 1 | 4 | 3 | 1 | 1
2017 | 42 | 8 | 5 | 3 | 11 | 5 | 4
2018 | 25 | 3 | 5 | 3 | 1 | 2
2019 | 40 | 9 | 2 | 28 | 1 | 2
Total: 259 issues

## Slide 44: Localized artefacts: 138 issues
94/138 local detector failure
27/138 originate from somewhere in the system or from the PMMA
  13/27 dust on PMMA test block
17/138 after equipment was serviced by technician

## Slide 45: Line problems: 28 issues
9/28 solved by calibration (service technician or local personnel)
12/28 no information
5/28 detector replacement
2/28 ?

## Slide 46: Defective blocks: 14 issues
7/14 detector replacement
4/14 calibration by service technician
2/14 no information
1/14 due to software upgrade, solved by intervention of service technician

## Slide 47: Ghost and lag problems: 38 issues
24/38 systematic ghosts, in a-Se based x-ray detectors
8/38 after service, solved by calibration (service technician or local personnel)
5/38 use of small compression paddle during calibration
1/38 ghost + line artefact + defective pixels

## Slide 48: Sudden changes in exposure signifiers: 18 issues
5/18 larger than normal variability
13/18 sudden changes after service or software upgrade
Mean pixelvalue
Snr

## Slide 49: Inhomogeneities: 15 issues
1/15 present during the whole life of the detector, unexplained
6/15 solved by calibration (service technician or local personnel)
1/15 grid problem
2/15 due to software upgrade, solved by new software version
1/15 Moiré pattern, solved by new software version
2/15 detector replacement
1/15 more than one issue (ghost + lines + pixels)
1/15 spontaneously resolved, unexplained

## Slide 50: Inhomogeneities: 15 issues

## Slide 51: Collimator problems: 7 issues
6/7 readjustment of collimator
  2/6 after planned service by technician
  2/6 visible in clinical images
  1/6 visible in clinical images, but not detected in DQC → classified as ‘Other’
1/7 inappropriate positioning of the PMMA block

## Slide 52: Results: trend analysis
105 systems: January 2015 (or when put into service) – January 2020
46/105 stable
54/105 sudden or gradual changes in exposure settings
  38/54 only sudden changes due to new software settings
  9/54 high variability
  7/54 sudden changes and gradual changes or high variability
5/105 steady change due to slow deterioration of component

## Slide 53: Limitations
Analysis has some subjective aspects
Analysis and alerting could be automated
Remedial action is not always communicated
Relevance of issues is not known
Noise power spectra are not analysed currently
Not proven that thumbnail images show all problems
Some changes go undetected e.g. sharpness, post processing

## Slide 54: Conclusion
Still large number of artefacts detected
Daily QC ensures constant quality of the system
Daily QC allows to find issues before they develop into major problems
In our service, we consider that the number of problems detected and solved justifies the effort of daily QC

## Slide 55: Ik hoop dat ik u later kan berichten uit Chicago…

## Slide 56: “A large retrospective study of non-compliance findings at half yearly medical physics tests of digital mammography devices using a vendor neutral (European) test protocol”
Door: Hannelore Verhoeven, Joke Binst, Kim Lemmens, Annelies Jacobs, Hilde Bosmans
Start: 1752 verslagen, van 131 toestellen, waarvan er 103 nog in dienst waren op het einde van 2020 en 28 hun volledige life cycle hebben gehad.
Submitted to RSNA 2021

## Slide 57: Typische problemen bij halfjaarlijkse testen
In totaal 472 problemen, waarvan 144 onmiddellijke actie vroegen.
homogeniteit 119 (no severe remarks),
alignment van het x-stralenveld 81 (80 severe),
compressie 69 (20 severe), dosis-dikte afregeling 42 (35 severe),
korte termijn reproduceerbaarheid 39, tube voltage 23 (1 severe),
verandering van de settings 29, dosimetrie 23,
local dense area test 14,
contrast detail 11 (8 severe),
software problems 7, noise 3 en andere 12.
Submitted to RSNA 2021

## Slide 58: 28 systemen over hun life cycle

## Slide 59: Basisprincipes van tomosynthese
Sechopoulos Med. Phys. 2013
New acq mode low dose 2D projections reconstructed vol

## Slide 60: Basisprincipes van tomosynthese
Courtesy Niklason, Radiology, 1997

## Slide 61: Verschillende designs mogelijk
Detectortechnologie
RX buis en anode/filter
  Uitdaging: een RX buis voor mammografie, met voldoende tube output (mGy/s)
Geen rooster of wel een rooster
  uitdaging: produktie en montage van een rooster is moeilijk

## Slide 62: Verschillende designs mogelijk
Balans tussen aantal projecties en dosis per projectie
Angular range; grotere hoek is betere 3D informatie
Dosisdistributie over de verschillende projecties
  Dosis = aantal projecties * dosis per projectie
  Uitdaging: elke projectie is een heel laag dosis beeld (ruis)
Dosis niveau
  uitdaging: hoe bepalen we dit? Hoe testen we de performantie?
‘Step and shoot’ versus ‘continue beweging van de focus’ of tussenoplossingen

## Slide 63: Verschillende designs mogelijk
Type reconstructie
  FBP of iteratieve reconstructie
  uitdaging: reconstructie tijd & test/optimaliseer performantie voor een grote reeks parameters
Reconstructie interval (bv 1mm coupes, of ‘slabbing’)
Link tussen het 2D beeld en tomosynthese (bv op display)

## Slide 64: Technische implementatie
Courtesy L Van Coillie

## Slide 65: Dosimetrie in Digitale Mammografie en DBT
Courtesy L Van Coillie

## Slide 66
Startend van de DBT stack, of van alle projecties, of van een deel van de projecties, berekenen van synthetisch ‘samenvattend beeld’
  mbv Maximum Intensity Projection (MIP), uitmiddeling, CAD technieken
  Andere post processing technieken om bv de visualisatie van de calcificaties te verbeteren
Deze synthetisch (berekende beelden) kunnen stand alone getoond worden of in combinatie met de tomo stack of andere beelden
New
Basisprincipes van synthetic 2D (SM)

## Slide 67
DM
DBT

## Slide 68
DBT
SM

## Slide 69: Het synthetisch 2D beeld is GEEN
2
… ander zicht, berekend van de tomosynthese stack (bv, het is GEEN CC berekend van een MLO acquisitie)
… eenvoudig iets; het is eigenlijk een heel gesofisticeerd beeld, dat in de toekomst nog beter, misschien ‘artificiëler’ zal worden met AI, etc. etc.

## Slide 70: De potentiële rol van het synthetisch 2D beeld
2
Snel overzicht van de borst
Side-by-side vergelijking van de borsten
Vergelijken met prior mammograms (2D FFDM/SM)
Helpt oriënteren en evalueren van microcalcificatieclusters
Dosis sparend potentieel (geen dubbele exposie van 2D FFDM + DBT)
-> maar hoe goed is het, voor massa’s en calcificaties?

## Slide 71: Hoe goed is het ? -> Clinical trials met DBT
Largely accomplished
Several US studies
STORM (Italy)
Oslo trial (Norway)
Malmo trial (Sweden)
Tommy trial (DBT only in recalled cases)
To start:
T-MIST (US)
TOSYMA (DE)
Tomosynthesis in the French screening, as soon as they have fixed their QC protocols

## Slide 72: An American study
JAMA. 2014 Jun 25;311(24):2499-507. doi: 10.1001/jama.2014.6095.
Breast cancer screening using tomosynthesis in combination with digital mammography.
Friedewald SM1, Rafferty EA2, Rose SL3, Durand MA4, Plecha DM5, Greenberg JS6, Hayes MK7, Copit DS8, Carlson KL9, Cink TM10, Barke LD11, Greer LN12, Miller DP13, Conant EF
for cancer detection, DM: 4.2 (95% CI, 3.8-4.7) vs DBT+DM: 5.4 (95% CI, 4.9-6.0) P < .001;
for invasive cancer detection, DM: 2.9 (95% CI, 2.5-3.2) vs DM+DBT: 4.1 (95% CI, 3.7-4.5) P < .001;
the in situ cancer detection rate was 1.4 (95% CI, 1.2-1.6) per 1000 screens with both methods.
adding tomosynthesis was associated with an increase in the positive predictive value for recall from 4.3% to 6.4%, P < .001

## Slide 73: A Swedish study
Eur Radiol. 2016 Jan;26(1):184-90. doi: 10.1007/s00330-015-3803-3. Epub 2015 May 1.
Performance of one-view breast tomosynthesis as a stand-alone breast cancer screeningmodality: results from the Malmö Breast Tomosynthesis Screening Trial, a population-based study.
Lång K1, Andersson I2, Rosso A3, Tingberg A4, Timberg P4, Zackrisson S2.
46 cases were detected by both modalities, 21 by DBT alone and 1 by DM alone.
Cancer detection rate for one-view DBT was 8.9/1000 screens (95 % CI 6.9 to 11.3) and 6.3/1000 screens (4.6 to 8.3) for two-view DM (p < 0.0001).
Recall rate after arbitration was 3.8 % (3.3 to 4.2) for DBT and 2.6 % (2.3 to 3.0) for DM (p < 0.0001).
PPV was 24 % for both DBT and DM.

## Slide 74: Toch nog twijfel over DBT voor screening
Nog te bewijzen: impact op interval cancer rate en impact op ….

## Slide 75: En hoe kan DBT geïmplementeerd worden?
Acquisitie & dosis setting
  Zijn alle systemen goed genoeg?
  1 view – 2 view
  FFDM and/or DBT
Data processing
  Reconstructie algoritme
  Synthetische mammografie
  CAD
  Deep learning technieken?
  Densiteit ook schatten?
Reading
  Slabs or slices
  Work station
  Second reading

## Slide 76: Quality Control

## Slide 77: Inhoudstafel van het EUREF QC protocol
Focal spot motion, tube output & tube voltage
Stralingsveld
Gemist weefsel aan thoraxzijde
Detector
  MTF, noise power
het complete systeem
  MTF voor verschillende hoogten in de borst
  CDMAM analyse voor follow up
Dosimetrie
Dosis settings als funktie van borstdikte
Ontbreekt:
  een limiet voor de performantie
  Een dagelijkse QC procedure

## Slide 78: Stand van zaken rond QC in DBT
Wij passen het EUREF protocol toe, en hebben zelf een performantietest gemaakt
EUREF document -> EFOMP document (work in progress)

## Slide 79: Quality Control
Nothing for synthetic 2D yet

## Slide 80: Ons testprotocol voor DBT en synthetic 2D
Ons ‘L1’ testobject (Leuven, by Lesley)
  Container in borstvorm, gevuld met groot aantal acrylic balletjes & water
  3D printed modellen van kanker
  Kleine calcs
CDMAM
“We present you our ECR 2018 presentation…”

## Slide 81: CONTRAST-DETAIL PHANTOM (CDMAM)
5
Aluminum plate with matrix of gold discs
  Disc thicknesses and diameters range
CDMAM placed within stack of 2 PMMA plates of 2 cm thick
  Equal to 5 cm PMMA or 6 cm breast (50% density)
Homogeneous background

## Slide 82
2D FFDM
½ AEC
AEC
2 AEC
DBT
SM
Notes: Higher the dose, less noisy

## Slide 83: TARGET DETECTABITLITY: CDMAM PHANTOM
Manual readout of CDMAM images
  Sara² software*
    Randomly shows and rotates each cell of CDMAM matrix
  5 readers
  10 images
    2D FFDM, DBT and SM
    ½ AEC, AEC and 2 AEC
  Nearest neighbor correction and polynomial curve fit
6
* Jacobs et al. A novel platform to simplify human observer performance experiments in clinical reading environments. SPIE Medical Imaging 2011

## Slide 84
CDMAM RESULTS
2D FFDM
DBT
SM

## Slide 85
CDMAM RESULTS
2D FFDM
DBT
SM
2D FFDM > DBT > SM (p<0.0001)
8

## Slide 86: STRUCTERED PHANTOM (L1)°
9
Semi-cylindrical container of PMMA
Background
  Equal volumes of PMMA spheres of different sizes
  Water
Clinically relevant lesions
  Microcalcifications: CaCO3 particles (Leeds test objects, UK)
  Spiculated and non-spiculated masses*: 3D printed
° Cockmartin et al. A structured phantom for comparative detectability evaluation in 2D digital mammography and breast tomosynthesis. ECR 2015
° Cockmartin et al. Comparison of detection performance between 2D digital mammography and breast tomosynthesis using a structured physical phantom. RSNA 2015
* Shaheen et al. The simulation of 3D mass models in 2D digital mammography and breast tomosynthesis. Med. Phys. 2014
Notes: In our research group from the KULeuven, a phantom was developed and launched. The phantom consists of a polymethyl methacrylate semi-cylindrical container and resembles a breast of 60 mm. The container is filled with water and PMMA spheres of different sizes. The space between the spheres was filled with water as this has x-ray properties close to breast fibrous tissue. The inserted lesions are microcalcifications and two different kind of masses. The microcalcifications exist out of 5 particles of calcium carbonate ordered like on a dice. The non-spiculated and spiculated masses were 3D printed This phantom is already tested for 2D FFDM and DBT for 5 different systems.

## Slide 87
2D FFDM
DBT
SM
Microcalcifications (180 - 200 µm)
½ AEC
AEC
2 AEC
Notes: Higher the dose, less noisy

## Slide 88
½ AEC
AEC
2 AEC
2D FFDM
DBT
SM
Non-spiculated mass (4.6 mm)
Notes: Higher the dose, less noisy

## Slide 89
½ AEC
AEC
2 AEC
2D FFDM
DBT
SM
Spiculated mass (9.7 mm)
Notes: Higher the dose, less noisy

## Slide 90: TARGET DETECTABITLITY: L1 PHANTOM
Manual readout of L1images
  In-house developed software for 4-AFC detection study
  5 readers
  10 images
    2D FFDM, DBT and SM
    ½ AEC, AEC and 2 AEC
    Microcalcifications, spiculated and non-spiculated masses
      5 sizes
  Average percent correctly detected targets (PC)
  Psychometric curve fit
    detection threshold diameter dtr at 62.5% correct respons
* Zhang et al. A four-alternative forced choice (4AFC) software for observer performance evaluation in radiology. SPIE Med. Imaging, Image Perception, Obs. Performance, Technol. Assess. 2016
10

## Slide 91
RESULTS L1
2D FFDM (dtr = 113.9 µm) > DBT (dtr = 145.6 µm) > SM (dtr = 151.5 µm)
Microcalcifications: AEC
11
Notes: For small diameters = ong microcalcs

## Slide 92
RESULTS L1
Microcalcifications: all doses
2D FFDM > DBT > SM
12
Notes: For small diameters = ong microcalcs

## Slide 93
RESULTS L1
Non-spiculated and spiculated masses: AEC
NS: DBT (dtr = 2.260 mm) > SM (dtr = 3.813 mm) > 2D FFDM (dtr = 3.435 mm)
S: DBT (dtr = 1.841 mm) > SM(dtr = 4.778 mm) > 2D FFDM (dtr = 4.839 mm)
13

## Slide 94
RESULTS L1
Non-spiculated and spiculated masses: all doses
NS: DBT > SM > 2D FFDM
S: DBT > SM > 2D FFDM
14

## Slide 95: CONCLUSION
For small diameter objects (~microcalcifications): 2D FFDM > DBT > SM
  CDMAM: 0.1 mm discs
  L1: microcalcifications
For larger diameter objects (~masses): DBT > SM > 2D FFDM
  L1: spiculated and non-spiculated masses
15

## Slide 96: Zal weldra gepubliceerd zijn:

## Slide 97: Take home points (Stay home points?)
Mammoscreening gaat gepaard met doorgedreven kwaliteitscontrole op alle niveaus.
  Beeldvorming, technische kwaliteit van het beeld, …
  Kankerdetectiegraad, detectie van kleine kankers, …
Een analyse na 15jaar dagelijkse kwaliteitscontrole toont dat er nog steeds nood is om deze inspanning verder te zetten.
Een analyse na 15 jaar acceptatie- en halfjaarlijkse testen toont dat ook deze inspanning moet verdergezet worden.

## Slide 98: Take home points (Stay home points?)
Het testprotocol voor DBT systemen is voldoende compleet
We kunnen ook synthetische mammo testen
Als we tomo in de screening wensen, moeten we kijken hoe het praktisch kan, en met een combinatie van technieken zodat de kwaliteit voor massa’s en microcalcificaties niet vermindert in vgl met 2D mammografie

## Slide 99: The role of DBT
For sure in diagnostic work up
In screening,
  if there is evidence that the screening performance will improve
  If work flow and costs are feasible
  If there is an international consensus
To depend on this international consensus, made me Google on
‘ongerust’

## Slide 100: Let’s get the pink mountain down, before the black one is taking over !